════════════════════════════════════════════════════════════════════════════════

# Risk Management File - ISO 14971:2019

════════════════════════════════════════════════════════════════════════════════

Company: InfuTech Medical Systems AG

Device: InfusoMax 3000 - Programmable Infusion Pump

Device Class: IIb

Generated: 2026-02-07

════════════════════════════════════════════════════════════════════════════════

*** SAMPLE DOCUMENT - FOR DEMONSTRATION ONLY ***

# RISK MANAGEMENT FILE

## InfusoMax 3000 Programmable Infusion Pump

### InfuTech Medical Systems AG

# PART 1: RISK MANAGEMENT PLAN

## 1.1 Scope and Purpose

### 1.1.1 Scope

### 1.1.2 Purpose

### 1.1.3 Device Overview

## 1.2 Risk Acceptability Criteria

### 1.2.1 Severity Classification Table

| Level | Severity | Description | Clinical Impact |
| --- | --- | --- | --- |
| S5 | Catastrophic | Death or permanent severe disability | Patient death, irreversible organ damage |
| S4 | Critical | Permanent moderate disability or temporary severe disability | Major injury requiring surgical intervention |
| S3 | Serious | Temporary moderate disability | Hospitalization extension, moderate injury |
| S2 | Minor | Temporary slight disability | Minor injury, temporary discomfort |
| S1 | Negligible | Inconvenience without injury | No clinical impact, minor inconvenience |

### 1.2.2 Probability Classification Table

| Level | Probability | Description | Frequency |
| --- | --- | --- | --- |
| P5 | Frequent | Very likely to occur repeatedly | > 1 in 100 |
| P4 | Probable | Likely to occur several times | 1 in 1,000 to 1 in 100 |
| P3 | Occasional | Likely to occur sometime | 1 in 10,000 to 1 in 1,000 |
| P2 | Remote | Unlikely to occur | 1 in 100,000 to 1 in 10,000 |
| P1 | Improbable | Very unlikely to occur | < 1 in 100,000 |

### 1.2.3 Risk Acceptability Matrix (5x5)

| Severity\Probability | P1 (Improbable) | P2 (Remote) | P3 (Occasional) | P4 (Probable) | P5 (Frequent) |
| --- | --- | --- | --- | --- | --- |
| S5 (Catastrophic) | ALARP | Unacceptable | Unacceptable | Unacceptable | Unacceptable |
| S4 (Critical) | Acceptable | ALARP | Unacceptable | Unacceptable | Unacceptable |
| S3 (Serious) | Acceptable | Acceptable | ALARP | Unacceptable | Unacceptable |
| S2 (Minor) | Acceptable | Acceptable | Acceptable | ALARP | ALARP |
| S1 (Negligible) | Acceptable | Acceptable | Acceptable | Acceptable | Acceptable |

### 1.2.4 Detectability Classification Table

| Level | Detectability | Description | Detection Capability |
| --- | --- | --- | --- |
| D1 | Very High | Almost certain detection | Automatic detection with immediate alert |
| D2 | High | High chance of detection | Detectable through routine monitoring |
| D3 | Moderate | Moderate chance of detection | Detectable through periodic checks |
| D4 | Low | Low chance of detection | Detection requires specific testing |
| D5 | Very Low | Very low chance of detection | Unlikely to be detected before harm |

## 1.3 Verification and Validation Activities

### 1.3.1 Risk Control Verification

### 1.3.2 Clinical Validation

## 1.4 Production and Post-Production Risk Management Activities

### 1.4.1 Production Risk Management

### 1.4.2 Post-Market Activities

## 1.5 Risk Management Team Roles and Responsibilities

| Role | Responsibility |
| --- | --- |
| Risk Management Lead | Overall risk management process oversight |
| Design Engineer | Technical risk assessment and control implementation |
| Clinical Affairs | Clinical risk evaluation and benefit assessment |
| Quality Assurance | Risk control verification and validation |
| Regulatory Affairs | Regulatory compliance and submission support |
| Post-Market Surveillance | Ongoing risk monitoring and assessment |

# PART 2: HAZARD IDENTIFICATION

## 2.1 Systematic Hazard Identification

| Hazard ID | Category | Hazard | Hazardous Situation | Harm | Initial Severity | Initial Probability | Initial Risk Level |
| --- | --- | --- | --- | --- | --- | --- | --- |
| H001 | Energy - Electrical | Electrical shock from power supply | User touches exposed electrical components during maintenance | Electric shock, burns, death | S5 | P2 | ALARP |
| H002 | Energy - Electrical | Power failure during infusion | Loss of power interrupts critical medication delivery | Therapeutic failure, patient deterioration | S4 | P3 | Unacceptable |
| H003 | Energy - Electrical | EMI interference | External electromagnetic fields disrupt pump operation | Incorrect infusion rate, therapy interruption | S3 | P3 | ALARP |
| H004 | Energy - Thermal | Overheating of internal components | Pump operates in high ambient temperature or ventilation blocked | Device malfunction, fire risk | S4 | P2 | ALARP |
| H005 | Energy - Mechanical | Mechanical failure of pump mechanism | Wear or damage to internal pumping components | Infusion stoppage or free flow | S5 | P2 | ALARP |
| H006 | Energy - Mechanical | Drop impact damage | Device dropped during transport or handling | Internal component damage, malfunction | S3 | P3 | ALARP |
| H007 | Biological | Contamination of fluid pathway | Breach in sterile fluid pathway during setup | Healthcare-associated infection | S4 | P2 | ALARP |
| H008 | Biological | Cross-contamination between patients | Inadequate cleaning/disinfection between uses | Healthcare-associated infection transmission | S4 | P2 | ALARP |
| H009 | Chemical | Drug incompatibility | Incompatible medications mixed in fluid pathway | Precipitation, therapeutic failure, toxicity | S4 | P3 | Unacceptable |
| H010 | Chemical | Chemical degradation of tubing | Certain medications degrade plastic tubing | Drug contamination, altered pharmacokinetics | S3 | P2 | Acceptable |
| H011 | Operational | Programming error - dose | User enters incorrect medication dose | Underdose or overdose of medication | S5 | P4 | Unacceptable |
| H012 | Operational | Programming error - rate | User enters incorrect infusion rate | Too fast or too slow medication delivery | S4 | P4 | Unacceptable |
| H013 | Operational | Incorrect drug selection | User selects wrong medication from drug library | Administration of wrong medication | S5 | P3 | Unacceptable |
| H014 | Operational | Failure to respond to alarms | User ignores or doesn't hear alarm notifications | Continued unsafe condition | S4 | P3 | Unacceptable |
| H015 | Operational | Improper IV line connection | IV tubing connected to wrong port or patient | Medication delivered to wrong patient | S5 | P3 | Unacceptable |
| H016 | Operational | Air-in-line not detected | Air detection system fails to identify air bubbles | Air embolism | S5 | P2 | ALARP |
| H017 | Operational | Occlusion not detected | Pressure sensing fails to detect line occlusion | Therapeutic failure, line rupture | S3 | P2 | Acceptable |
| H018 | Operational | Free flow of medication | Failure of flow control mechanism | Rapid delivery of entire fluid volume | S5 | P2 | ALARP |
| H019 | Operational | Inadequate user training | User lacks proper training on device operation | Multiple use errors, patient harm | S4 | P3 | Unacceptable |
| H020 | Information | Display malfunction | Screen becomes unreadable or shows incorrect data | User cannot monitor infusion status | S3 | P2 | Acceptable |
| H021 | Information | Incorrect alarm settings | Alarm limits set inappropriately for patient condition | Late detection of problems | S4 | P3 | Unacceptable |
| H022 | Information | Software malfunction | Software error causes incorrect operation | Various patient safety issues | S4 | P2 | ALARP |
| H023 | Information | Data logging failure | Infusion history not properly recorded | Inability to track therapy, dosing errors | S2 | P3 | Acceptable |
| H024 | Information | Network communication failure | Pump cannot communicate with hospital systems | Loss of remote monitoring capability | S2 | P3 | Acceptable |
| H025 | Environmental | Exposure to cleaning chemicals | Harsh disinfectants damage pump components | Device malfunction after cleaning | S3 | P3 | ALARP |
| H026 | Environmental | Operating outside temperature range | Device used in extreme temperature conditions | Performance degradation, malfunction | S3 | P2 | Acceptable |
| H027 | Environmental | Exposure to moisture/liquids | Spilled fluids enter pump housing | Electrical short, corrosion, malfunction | S4 | P3 | Unacceptable |
| H028 | Environmental | Physical impact/vibration | Excessive vibration or shock during transport | Component damage, calibration drift | S3 | P2 | Acceptable |
| H029 | Operational | Battery failure | Internal battery fails to maintain power during outage | Loss of therapy during power interruption | S4 | P3 | Unacceptable |
| H030 | Operational | Tubing misloading | IV tubing not properly loaded in pump mechanism | Incorrect flow rate, no flow | S4 | P4 | Unacceptable |

# PART 3: PROCESS FMEA (pFMEA)

## 3.1 Manufacturing Process Risk Analysis

| ID | Process Step | Potential Failure Mode | Effect | Severity | Cause | Probability | Detection | RPN | Risk Control | Residual Risk |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| pF001 | PCB Assembly | Incorrect component placement | Device malfunction, safety feature failure | 4 | Human error, programming error | 3 | 2 | 24 | Automated placement verification, visual inspection | ALARP |
| pF002 | PCB Assembly | Cold solder joint | Intermittent electrical connection | 4 | Insufficient heat, contamination | 2 | 3 | 24 | Process validation, automated optical inspection | Acceptable |
| pF003 | Software Installation | Wrong software version loaded | Incorrect device operation | 4 | Version control error | 2 | 2 | 16 | Automated version verification, checksum validation | Acceptable |
| pF004 | Pump Mechanism Assembly | Improper torque on fasteners | Mechanical failure during use | 4 | Torque wrench malfunction, human error | 2 | 2 | 16 | Torque verification, calibrated tools | Acceptable |
| pF005 | Pressure Sensor Calibration | Incorrect calibration values | False alarms or missed occlusions | 3 | Calibration equipment drift | 2 | 2 | 12 | Regular calibration verification, standards traceability | Acceptable |
| pF006 | Display Assembly | Screen adhesive failure | Display separation, readability issues | 3 | Adhesive degradation, contamination | 2 | 3 | 18 | Environmental testing, adhesive qualification | Acceptable |
| pF007 | Power Supply Testing | Inadequate electrical safety testing | Electrical hazard to user | 5 | Test equipment failure, procedure error | 2 | 2 | 20 | Redundant testing, equipment calibration | ALARP |
| pF008 | Housing Assembly | Inadequate seal integrity | Moisture ingress | 3 | Gasket damage, improper installation | 3 | 2 | 18 | Leak testing, gasket inspection | Acceptable |
| pF009 | Final Functional Test | Test not performed completely | Defective unit shipped | 4 | Time pressure, incomplete procedures | 3 | 2 | 24 | Automated test sequences, verification checkpoints | ALARP |
| pF010 | Alarm System Test | Alarm volume insufficient | User may not hear alarms | 4 | Speaker defect, test procedure gap | 2 | 3 | 24 | Acoustic testing, multiple test points | ALARP |
| pF011 | Battery Installation | Wrong battery type installed | Power system failure | 4 | Part number confusion, supplier error | 2 | 2 | 16 | Part number verification, incoming inspection | Acceptable |
| pF012 | Flow Sensor Assembly | Sensor misalignment | Inaccurate flow measurement | 4 | Fixturing error, dimensional variation | 2 | 2 | 16 | Alignment verification, functional testing | Acceptable |
| pF013 | Packaging | Damage during packaging | Product damage before use | 2 | Rough handling, inadequate protection | 3 | 3 | 18 | Packaging validation, handling procedures | Acceptable |
| pF014 | Quality Control | Sampling error in inspection | Defective units not detected | 4 | Statistical sampling error | 2 | 4 | 32 | Enhanced sampling plans, inspector training | ALARP |
| pF015 | Sterilization | Insufficient sterilization | Bioburden on device surfaces | 3 | Process parameter deviation | 2 | 2 | 12 | Process monitoring, biological indicators | Acceptable |

# PART 4: PRODUCT FMEA (dFMEA)

## 4.1 Design Risk Analysis

| ID | Component/Function | Potential Failure Mode | Effect | Severity | Cause | Probability | Detection | RPN | Risk Control | Residual Risk |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| dF001 | Pump Motor | Motor stall/failure | Infusion stops, no medication delivery | 4 | Mechanical wear, electrical failure | 2 | 1 | 8 | Motor current monitoring, redundant flow sensing | Acceptable |
| dF002 | Pressure Sensor | False occlusion alarm | Unnecessary therapy interruption | 3 | Sensor drift, contamination | 3 | 2 | 18 | Regular calibration, self-test function | Acceptable |
| dF003 | Air-in-Line Detector | Fails to detect air bubbles | Air embolism risk | 5 | Optical sensor contamination | 2 | 2 | 20 | Redundant detection methods, self-test | ALARP |
| dF004 | Flow Control Valve | Valve sticks open | Free flow of medication | 5 | Contamination, mechanical wear | 2 | 1 | 10 | Multiple valve design, flow monitoring | Acceptable |
| dF005 | Touchscreen Display | Display becomes unresponsive | Cannot control or monitor infusion | 4 | Software hang, hardware failure | 2 | 1 | 8 | Watchdog timer, hardware buttons backup | Acceptable |
| dF006 | Power Supply | Output voltage regulation failure | Device malfunction or shutdown | 4 | Component aging, thermal stress | 2 | 1 | 8 | Voltage monitoring, backup power | Acceptable |
| dF007 | Battery System | Battery does not charge | No backup power during outage | 4 | Charging circuit failure | 2 | 2 | 16 | Battery status monitoring, charge indicators | Acceptable |
| dF008 | CPU/Processor | Processor lockup | Complete device failure | 4 | Software bug, hardware failure | 2 | 1 | 8 | Watchdog timer, fault detection | Acceptable |
| dF009 | Memory (RAM) | Data corruption | Incorrect infusion parameters | 4 | Cosmic radiation, electrical noise | 2 | 2 | 16 | Error correction codes, data validation | Acceptable |
| dF010 | Memory (Flash) | Program corruption | Device inoperable | 4 | Write/erase cycles, power loss | 1 | 1 | 4 | Checksums, dual boot capability | Acceptable |
| dF011 | Alarm Speaker | Speaker failure | Audible alarms not heard | 4 | Component failure, connection loss | 2 | 3 | 24 | Redundant alarm methods, self-test | ALARP |
| dF012 | Tubing Sensor | Fails to detect tubing presence | Pump operates without tubing | 4 | Sensor failure, mechanical damage | 2 | 1 | 8 | Multiple sensing methods, interlock system | Acceptable |
| dF013 | Door/Access Panel | Door opens during operation | Tubing displacement, flow disruption | 3 | Latch failure, user error | 3 | 1 | 9 | Door sensor, automatic stop function | Acceptable |
| dF014 | Network Interface | Communication failure | Loss of remote monitoring | 2 | Hardware failure, network issues | 3 | 2 | 12 | Local operation capability, status indicators | Acceptable |
| dF015 | Drug Library | Corrupted drug data | Wrong dosing limits applied | 5 | Software error, data corruption | 1 | 2 | 10 | Data integrity checks, backup storage | Acceptable |
| dF016 | Flow Rate Sensor | Inaccurate flow measurement | Wrong medication delivery rate | 4 | Sensor drift, mechanical wear | 2 | 2 | 16 | Regular calibration, cross-check with other sensors | Acceptable |
| dF017 | Keypad/Buttons | Button does not respond | Cannot enter parameters | 3 | Mechanical wear, contamination | 3 | 1 | 9 | Touchscreen backup, tactile feedback | Acceptable |
| dF018 | Housing/Enclosure | Crack or breach in housing | Contamination, moisture ingress | 3 | Impact damage, material fatigue | 2 | 3 | 18 | Impact testing, material selection | Acceptable |
| dF019 | Mounting System | Mount fails during use | Device falls, physical damage | 3 | Mechanical failure, overload | 2 | 1 | 6 | Load testing, safety factor in design | Acceptable |
| dF020 | Real-Time Clock | Clock drift/failure | Incorrect time stamps, scheduling errors | 2 | Crystal aging, temperature variation | 3 | 3 | 18 | Network time sync, user notification | Acceptable |

# PART 5: CLINICAL FMEA

## 5.1 Clinical Use Risk Analysis

| ID | Clinical Scenario | Potential Failure Mode | Clinical Effect | Severity | Cause | Probability | Detection | RPN | Risk Control | Residual Risk |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| cF001 | Critical care medication infusion | Overdose due to programming error | Patient toxicity, organ damage | 5 | Nurse enters wrong dose, decimal point error | 3 | 2 | 30 | Drug library with dosing limits, double verification | ALARP |
| cF002 | Chemotherapy administration | Wrong medication selected | Inappropriate chemotherapy, delayed treatment | 4 | Similar drug names, user confusion | 3 | 2 | 24 | Distinctive drug library entries, barcode scanning | ALARP |
| cF003 | Pain management infusion | Underdose of analgesic | Inadequate pain control, patient suffering | 3 | Programming error, pump malfunction | 3 | 3 | 27 | Pain assessment protocols, backup pain management | ALARP |
| cF004 | Parenteral nutrition delivery | Interruption of nutrition therapy | Malnutrition, delayed recovery | 3 | Pump failure, tubing occlusion | 2 | 2 | 12 | Backup feeding methods, close monitoring | Acceptable |
| cF005 | Insulin infusion in ICU | Hypoglycemic event from overdose | Coma, brain damage, death | 5 | Calculation error, pump malfunction | 2 | 2 | 20 | Glucose monitoring protocols, smart dosing algorithms | ALARP |
| cF006 | Antibiotic therapy | Subtherapeutic antibiotic levels | Treatment failure, resistance development | 3 | Underdosing, infusion interruption | 3 | 3 | 27 | Therapeutic monitoring, backup administration routes | ALARP |
| cF007 | Fluid resuscitation | Inadequate fluid delivery | Hemodynamic instability, shock | 4 | Flow rate error, line disconnection | 2 | 2 | 16 | Hemodynamic monitoring, multiple access sites | Acceptable |
| cF008 | Pediatric medication infusion | Overdose in small patient | Enhanced drug sensitivity, toxicity | 5 | Weight-based calculation error | 2 | 2 | 20 | Pediatric drug library, weight verification | ALARP |
| cF009 | Sedation during procedures | Oversedation | Respiratory depression, airway compromise | 4 | Incorrect dosing, drug accumulation | 2 | 2 | 16 | Continuous monitoring, reversal agents available | Acceptable |
| cF010 | Emergency medication administration | Delayed drug delivery | Therapeutic failure in critical situation | 4 | Complex programming, user error under stress | 3 | 2 | 24 | Emergency drug protocols, simplified interfaces | ALARP |

# PART 6: USABILITY FMEA

## 6.1 Use-Related Risk Analysis

| ID | Use Task | Use Error / Misuse | Hazardous Situation | Harm | Severity | Probability | Detection | RPN | Risk Control | Residual Risk |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| uF001 | Programming drug dose | Entering dose in wrong units (mg vs mcg) | 1000x overdose administered | Drug toxicity, death | 5 | 3 | 2 | 30 | Unit selection prominent, confirmation screens | ALARP |
| uF002 | Setting infusion rate | Confusion between rate and volume | Too rapid or too slow delivery | Therapeutic failure, toxicity | 4 | 3 | 2 | 24 | Clear labels, separate input fields | ALARP |
| uF003 | Connecting IV tubing | Tubing connected backwards | No drug delivery | Therapeutic failure | 4 | 2 | 1 | 8 | Keyed connectors, visual indicators | Acceptable |
| uF004 | Responding to alarms | Silencing alarm without addressing cause | Continued unsafe condition | Patient harm | 4 | 3 | 3 | 36 | Escalating alarms, mandatory acknowledgment | ALARP |
| uF005 | Loading tubing in pump | Tubing not fully seated in mechanism | Inaccurate flow rate | Wrong drug delivery | 3 | 3 | 1 | 9 | Tubing presence sensors, load verification | Acceptable |
| uF006 | Selecting from drug library | Selecting similar-named drug | Wrong medication administered | Adverse drug reaction | 5 | 2 | 2 | 20 | Distinctive drug names, confirmation prompts | ALARP |
| uF007 | Setting alarm limits | Setting limits too wide | Late detection of problems | Delayed intervention | 3 | 3 | 3 | 27 | Default limits, range restrictions | ALARP |
| uF008 | Reading display information | Misreading small text/numbers | Incorrect interpretation of status | Inappropriate clinical decisions | 3 | 3 | 3 | 27 | Large fonts, high contrast display | ALARP |
| uF009 | Cleaning/disinfecting device | Using inappropriate cleaning agent | Device damage, contamination | Device failure, infection | 4 | 2 | 2 | 16 | Cleaning instructions, material compatibility | Acceptable |
| uF010 | Battery management | Not connecting to power during use | Battery depletion, therapy interruption | Therapeutic failure | 3 | 3 | 1 | 9 | Battery status indicators, low power alarms | Acceptable |

# PART 7: RISK CONTROL MEASURES

## 7.1 Risk Control Options Analysis

### 7.1.1 Inherent Safety by Design

### 7.1.2 Protective Measures in Device

### 7.1.3 Information for Safety

## 7.2 Risk Control Measures Summary

| Risk ID | Hazard | Risk Control Measure | Type | Verification Method | Residual Risk | Acceptable? |
| --- | --- | --- | --- | --- | --- | --- |
| H001 | Electrical shock | Double insulation, protective earthing | Inherent Design | IEC 60601-1 testing | Acceptable | Yes |
| H002 | Power failure | Battery backup system, power loss alarms | Protective Measure | Battery life testing | Acceptable | Yes |
| H003 | EMI interference | EMC design, shielding | Inherent Design | IEC 60601-1-2 testing | Acceptable | Yes |
| H004 | Overheating | Temperature monitoring, thermal shutdown | Protective Measure | Thermal testing | Acceptable | Yes |
| H005 | Mechanical failure | Multiple flow sensors, pump monitoring | Protective Measure | Reliability testing | Acceptable | Yes |
| H006 | Drop impact | Robust housing design, shock mounting | Inherent Design | Drop testing per IEC 60601-1 | Acceptable | Yes |
| H007 | Fluid pathway contamination | Sterile disposable sets, closed system | Inherent Design | Sterility testing | Acceptable | Yes |
| H008 | Cross-contamination | Cleaning procedures, non-contact surfaces | Information | Cleaning validation | Acceptable | Yes |
| H009 | Drug incompatibility | Drug library compatibility checks | Protective Measure | Clinical validation | Acceptable | Yes |
| H010 | Chemical tubing degradation | Compatible materials, usage guidelines | Inherent Design | Material testing | Acceptable | Yes |
| H011 | Dose programming error | Drug library dose limits, confirmation | Protective Measure | Usability testing | Acceptable | Yes |
| H012 | Rate programming error | Rate limit checks, unit verification | Protective Measure | Software testing | Acceptable | Yes |
| H013 | Wrong drug selection | Distinctive naming, barcode integration | Protective Measure | Usability validation | Acceptable | Yes |
| H014 | Ignoring alarms | Escalating alarms, visual indicators | Protective Measure | Alarm testing | Acceptable | Yes |
| H015 | Wrong line connection | Color coding, unique connectors | Inherent Design | Human factors testing | Acceptable | Yes |
| H016 | Air detection failure | Redundant air sensors, self-test | Protective Measure | Sensitivity testing | Acceptable | Yes |
| H017 | Occlusion detection failure | Multiple pressure sensors, trending | Protective Measure | Pressure testing | Acceptable | Yes |
| H018 | Free flow | Anti-free-flow valve, flow monitoring | Inherent Design | Flow testing | Acceptable | Yes |
| H019 | Inadequate training | Comprehensive training program, competency testing | Information | Training validation | Acceptable | Yes |
| H020 | Display malfunction | Display self-test, backup indicators | Protective Measure | Display testing | Acceptable | Yes |
| H021 | Wrong alarm settings | Default limits, range restrictions | Protective Measure | Software validation | Acceptable | Yes |
| H022 | Software malfunction | Software testing, fault handling | Inherent Design | IEC 62304 compliance | Acceptable | Yes |
| H023 | Data logging failure | Redundant storage, error detection | Protective Measure | Data integrity testing | Acceptable | Yes |
| H024 | Network failure | Local operation capability, status display | Inherent Design | Network testing | Acceptable | Yes |
| H025 | Chemical exposure | Chemical-resistant materials, IP rating | Inherent Design | Chemical compatibility testing | Acceptable | Yes |
| H026 | Temperature extremes | Operating range specification, warnings | Information | Environmental testing | Acceptable | Yes |
| H027 | Moisture exposure | IP54 rating, sealed enclosure | Inherent Design | Ingress protection testing | Acceptable | Yes |
| H028 | Physical impact | Shock-resistant design, mounting | Inherent Design | Vibration/shock testing | Acceptable | Yes |
| H029 | Battery failure | Battery monitoring, replacement indicators | Protective Measure | Battery testing | Acceptable | Yes |
| H030 | Tubing misloading | Tubing guides, presence sensors | Inherent Design | Loading mechanism testing | Acceptable | Yes |

# PART 8: RISK-BENEFIT ANALYSIS

## 8.1 Overall Residual Risk Evaluation

## 8.2 Risk-Benefit Determination

### 8.2.1 Clinical Benefits

### 8.2.2 Risk-Benefit Assessment

## 8.3 Comparison with State of the Art

## 8.4 Conclusions

| Residual Risk Category | Count | Percentage | Acceptability |
| --- | --- | --- | --- |
| Acceptable | 27 | 90.0% | Within acceptable limits |
| ALARP | 3 | 10.0% | As Low As Reasonably Practicable |
| Unacceptable | 0 | 0.0% | None remaining |
| Total Risks | 30 | 100.0% | All controlled |

# PART 9: RISK MANAGEMENT REPORT

## 9.1 Summary of Risk Management Activities

## 9.2 Risk Management Plan Fulfillment

### 9.2.1 Completed Activities

### 9.2.2 Verification Activities

## 9.3 Overall Residual Risk Acceptability Statement

## 9.4 Production and Post-Production Information Review

### 9.4.1 Production Risk Management

### 9.4.2 Post-Market Surveillance Plan

## 9.5 Conclusions and Approval

### 9.5.1 Risk Management Conclusions

### 9.5.2 Approval

════════════════════════════════════════════════════════════════════════════════

# APPROVAL SIGNATURES

Prepared by: _________________ Date: __________

Reviewed by: _________________ Date: __________

Approved by: _________________ Date: __________

*** SAMPLE DOCUMENT - FOR DEMONSTRATION ONLY ***